21 CFR PART 11 COMPLIANCE VERIFICATION CHECKLIST

Tap Density Tester — Model TD-2B

| Document No. | ________________________________ |
| --- | --- |
| Revision No. | 02 (Supersedes Rev. 01) |
| Equipment / Instrument Name | Tap Density Tester |
| Model No. | TD-2B |
| Equipment ID / Asset No. | ________________________________ |
| Location / Department | Quality Control Laboratory |
| Software Version | ________________________________ |
| Checklist Reference SOP No. | ________________________________ |
| Date of Execution | ________________________________ |

Revision 02 Note: This revision expands the original checklist to add Power Failure handling and auto-approval (Section 7), USB1/USB2 port validation (Section 9), SP/Recipe creation with mandatory segregation of duties — creator cannot self-approve (Section 4), mandatory password change on first login (Section 3), and historical test-report reprint verification of batch number, product name, and start/end time (Section 10). These additions were incorporated after cross-referencing the Installation & Operational Qualification protocols for the RFT-2B Tablet Friability Tester and RHT-3PB Tablet Hardness Tester.

1. Purpose

To verify, through documented objective evidence, that the electronic records and electronic signature features of the Tap Density Tester (Model TD-2B) comply with the applicable requirements of 21 CFR Part 11, and that its system security, individual function-level access control, segregation of duties, audit trail, electronic approval workflow, power-failure resilience, and record integrity all perform as intended.

2. Scope

This checklist applies to the Tap Density Tester, Model TD-2B, and covers verification of system/master settings, user creation and individual function-level access rights, password and first-login security controls, SP/Recipe creation with segregation-of-duties approval, test execution (Quick Test and SP/Recipe Test), electronic approval workflow (including biometric and password-based signature methods), power-failure handling, date-and-time security, calibration and USB port validation, reporting and historical record integrity, audit trail generation, and login/session security controls. This checklist addresses data-integrity / 21 CFR Part 11 verification only; analytical or metrological performance qualification of the instrument (e.g., tap count accuracy, drop height, cylinder volume) shall be addressed separately under the equipment Performance Qualification (PQ) protocol.

3. Instructions

- Perform each check point sequentially and record the actual observation in the "Remarks" column.

- Mark the result as Pass or Fail for each test case.

- Attach supporting objective evidence (Audit Trail printout/screenshot, test report, etc.) for each section and cross-reference the attachment number in the "Remarks" column.

- Where a check requires two distinct users (e.g., SP/Recipe segregation of duties in Section 4, or Profile Management edits in Sections 1–2), use two separate User ID logins and confirm both are correctly reflected in the Audit Trail.

- Any "Fail" / non-compliant observation shall be recorded as a deviation and routed as per the site Deviation Management SOP prior to closure of this checklist.

4. Checklist

| Test Case ID | Test Case Description | User Role | Acceptance Criteria | Pass / Fail | Remarks |
| --- | --- | --- | --- | --- | --- |
| 1. USER MANAGEMENT & ROLE CONTROL [Ref: 21 CFR 11.10(d), 11.10(g)] |  |  |  |  |  |
| OQ-UM-01 | Administrator Creation | Administrator | System allows creation of a new Administrator with a unique User ID and password. |  |  |
| OQ-UM-02 | Administrator Disabling | Administrator | Disabled Administrator account cannot log in. |  |  |
| OQ-UM-03 | Reviewer Creation | Administrator | System allows creation of a Reviewer account with selectable privileges. |  |  |
| OQ-UM-04 | Reviewer Disabling | Administrator | Disabled Reviewer cannot log in or perform any operation. |  |  |
| OQ-UM-05 | User Creation | Administrator | System allows creation of a User (Operator) account with selectable privileges. |  |  |
| OQ-UM-06 | User Disabling | Administrator | Disabled User cannot log in or access any function. |  |  |
| OQ-UM-07 | QA Creation | Administrator | System allows creation of a QA account with selectable privileges. |  |  |
| OQ-UM-08 | QA Disabling | Administrator | Disabled QA cannot log in or access any function. |  |  |
| OQ-UM-09 | User Profile Edit Restriction | Any user without Profile Management right | A user who does not hold Profile Management access cannot create, edit, or disable any user account; the option is unavailable or blocked. |  |  |
| OQ-UM-AT | Audit Trail Check — User Management | Reviewer/QA | All user creation, disabling, and editing actions — including edits made by whoever holds Profile Management access — appear in the Audit Trail with user ID, date & time. |  |  |
| 2. FUNCTION-LEVEL PERMISSION (INDIVIDUAL POWER) CONFIGURATION [Ref: 21 CFR 11.10(d), 11.10(g), 11.200] |  |  |  |  |  |
| OQ-RP-01 | Individual Function Assignment | Administrator / Profile Management holder | Each individual system function (e.g., Test Functionality, Recipe/SP Management, Recipe/SP Approval, User Management, Profile Management, Real-Time Clock Control, Export Reports, Audit Trail access, Calibration, and other applicable functions) can be independently enabled or disabled per user at the time of user creation or edit. |  |  |
| OQ-RP-02 | Assignment Restricted to Authorised Role | Any user | Only a user holding Administrator rights or Profile Management access can decide or change which functions are assigned to any user; this option is unavailable to all other users. |  |  |
| OQ-RP-03 | Individual Power Enforcement | Assigned User | Logging in as a user with only specific functions assigned, verify that ONLY those assigned functions are accessible and operate correctly, and every non-assigned function is inaccessible or blocked. |  |  |
| OQ-RP-AT | Audit Trail Check — Permission Configuration | Reviewer/QA | Every function/permission assigned, changed, or removed for a user is recorded in the Audit Trail with user ID, date & time. |  |  |
| 3. PASSWORD & SECURITY CONTROLS [Ref: 21 CFR 11.10(d), 11.300] |  |  |  |  |  |
| OQ-SEC-01 | Password Change — Administrator | Administrator | Administrator can change own password successfully. |  |  |
| OQ-SEC-02 | Password Change — Reviewer | Reviewer | Reviewer can change own password successfully. |  |  |
| OQ-SEC-03 | Password Change — User | User | User can change own password successfully. |  |  |
| OQ-SEC-04 | Password Change — QA | QA | QA can change own password successfully. |  |  |
| OQ-SEC-05 | Mandatory Password Change on First Login | Newly created User (any role) | On first login after creation, the system mandatorily prompts for a password change before allowing access to the home screen. The screen requires entry of the current/temporary password, the new password, and re-entry of the new password for confirmation; login is blocked until this is completed successfully. |  |  |
| OQ-SEC-06 | Multiple Wrong Password Attempts | Any User | Account locks automatically after three (3) consecutive failed login attempts (or the configured limit). |  |  |
| OQ-SEC-07 | Account Unlocking | Administrator | Only an Administrator can unlock a locked user account. |  |  |
| OQ-SEC-08 | Password Policy Enforcement | All Roles | System enforces the defined password rules (length, complexity, expiry, as applicable). |  |  |
| OQ-SEC-AT | Audit Trail Check — Security | Reviewer/QA | Password changes, first-login password changes, failed login attempts, lockouts, and unlocks are all recorded in the Audit Trail. |  |  |
| 4. RECIPE / SP (TEST METHOD) CREATION — SEGREGATION OF DUTIES & APPROVAL [Ref: 21 CFR 11.10(g), 11.50, 11.70] |  |  |  |  |  |
| OQ-RC-01 | SP / Recipe Creation | Authorised User / Administrator | System allows creation of a Recipe / SP (test method) defining product name, test parameters, and acceptance limits. |  |  |
| OQ-RC-02 | SP / Recipe Pending-Approval State | System | A newly created or edited SP/Recipe is held in a "Pending Approval" status and cannot be selected for testing until approved. |  |  |
| OQ-RC-03 | Segregation of Duties — Creator Cannot Self-Approve | Creator of the SP/Recipe | The user who created (or last edited) the SP/Recipe cannot also approve it; if the same User ID attempts approval, the system denies the action. |  |  |
| OQ-RC-04 | SP / Recipe Approval | User holding SP Approval right | A different, authorised user holding SP/Recipe approval rights can approve the SP/Recipe; only an approved SP/Recipe becomes available for testing. |  |  |
| OQ-RC-05 | SP / Recipe Rejection | User holding SP Approval right | Approver can reject an SP/Recipe; a rejected SP/Recipe cannot be used for testing. |  |  |
| OQ-RC-06 | SP / Recipe Edit Restriction | Any user | If an approved SP/Recipe is edited, it automatically reverts to "Pending Approval" and requires fresh approval before reuse. |  |  |
| OQ-RC-AT | Audit Trail Check — SP/Recipe | Reviewer/QA | Creation, edit, approval, and rejection of every SP/Recipe are recorded in the Audit Trail, capturing both the creator's User ID and the approver's User ID, with date & time. |  |  |
| 5. TEST EXECUTION — QUICK TEST & SP (RECIPE) TEST [Ref: 21 CFR 11.10(a), 11.10(f)] |  |  |  |  |  |
| OQ-TE-01 | Quick Test Execution | User | Quick Test runs correctly from start to result generation without error. |  |  |
| OQ-TE-02 | SP / Recipe Test Execution | User | Test executes correctly using an approved SP/Recipe, from start to result generation, without error. |  |  |
| OQ-TE-AT | Audit Trail Check — Test Execution | Reviewer | Test start, parameters used, and result generation are recorded in the Audit Trail with user ID, date & time. |  |  |
| 6. TEST APPROVAL WORKFLOW & ELECTRONIC SIGNATURE [Ref: 21 CFR 11.50, 11.70, 11.200] |  |  |  |  |  |
| OQ-WF-01 | Pre-Approval Printout / Preview | User / Reviewer | After a test completes and before approval, the system displays or prints a preview clearly marked as "Not Approved" (draft status). |  |  |
| OQ-WF-02 | Approval Method — User ID & Password | Reviewer | Test approval can be completed successfully using User ID and password entry. |  |  |
| OQ-WF-03 | Approval Method — Biometric / Fingerprint | Reviewer | Test approval can be completed successfully using fingerprint / biometric authentication, where enabled. |  |  |
| OQ-WF-04 | Approval of a PASS Result | Reviewer | A completed test with a PASS result can be reviewed and approved correctly. |  |  |
| OQ-WF-05 | Approval of a FAIL Result | Reviewer | A completed test with a FAIL result can be reviewed and approved correctly; the approval workflow functions identically for a Fail outcome. |  |  |
| OQ-WF-06 | Final Report Generated Only Post-Approval | System | The final, approved test report is generated only after the approval step is completed; no approved report is available beforehand. |  |  |
| OQ-WF-AT | Audit Trail Check — Approval | Reviewer/QA | The approval action — approver User ID, authentication method used, Pass/Fail outcome, date & time — is recorded in the Audit Trail for every test. |  |  |
| 7. POWER FAILURE HANDLING [Ref: 21 CFR 11.10(a), 11.10(e)] |  |  |  |  |  |
| OQ-PF-01 | Power Interruption During an Active Test | User | With a test in progress, switch OFF the instrument at the rear / mains power switch; verify the instrument loses power immediately. |  |  |
| OQ-PF-02 | Power Restoration & Status on Re-Login | User | Power ON the instrument and log back in; verify the system flags / displays the interrupted test with a "Power Failure" status. |  |  |
| OQ-PF-03 | Auto-Abort of Interrupted Test | System | The interrupted test record is automatically marked "Aborted – Power Failure." |  |  |
| OQ-PF-04 | Auto-Save & Auto-Approval on Power Failure | System | The aborted record is automatically saved and auto-approved by the system, with the remark/reason recorded as "Auto-Approved – Power Failure." |  |  |
| OQ-PF-AT | Audit Trail Check — Power Failure | Reviewer/QA | The power-failure event, auto-abort, and auto-approval are all recorded in the Audit Trail with date & time. |  |  |
| 8. DATE & TIME CONTROLS [Ref: 21 CFR 11.10(a), 11.10(e)] |  |  |  |  |  |
| OQ-SYS-01 | Real-Time Clock (RTC) Setting | Administrator | Administrator can set / change the system date & time correctly. |  |  |
| OQ-SYS-02 | Date/Time Retention After Power Cycle | Administrator | After configuring date/time, power OFF and ON the instrument; the previously set date/time is retained accurately and not reverted. |  |  |
| OQ-SYS-AT | Audit Trail Check — Date/Time Edit | Reviewer/QA | Every date/time change — including the specific old value and new value — is recorded in the Audit Trail with the User ID and timestamp of the edit itself. |  |  |
| 9. CALIBRATION & VALIDATION (INCL. USB1 / USB2 VALIDATION) [Ref: 21 CFR 11.10(a), 11.10(h)] |  |  |  |  |  |
| OQ-CV-01 | Instrument Calibration | Administrator | The instrument's measurement parameter(s) are calibrated / verified against the designated certified reference standard(s) specified in the operating manual / site calibration SOP, with the displayed value within acceptance limits. |  |  |
| OQ-CV-02 | Instrument Validation | Administrator | System accuracy is validated using certified reference standards traceable to recognised national standards; results are within acceptance limits. |  |  |
| OQ-CV-03 | Calibration / Validation Restricted to Administrator | User | User cannot perform or alter calibration/validation; the function is restricted to Administrator (or a user explicitly granted this right). |  |  |
| OQ-CV-04 | USB1 Port Validation | Administrator | The USB1 port validation function completes successfully and correctly confirms the port / connected-device status. |  |  |
| OQ-CV-05 | USB2 Port Validation | Administrator | The USB2 port validation function completes successfully and correctly confirms the port / connected-device status. |  |  |
| OQ-CV-AT | Audit Trail & Report Check — Calibration/Validation/USB | Reviewer/QA | Calibration, validation, and USB1/USB2 validation events are all recorded in the Audit Trail with user ID, date & time; a corresponding validation report/printout is generated for each. |  |  |
| 10. REPORTING & DATA INTEGRITY [Ref: 21 CFR 11.10(b), 11.10(c)] |  |  |  |  |  |
| OQ-RPT-01 | Report Generation | User | System generates a test report successfully after test approval. |  |  |
| OQ-RPT-02 | Thermal Printer Output | User | Inbuilt thermal printer prints the report clearly and completely. |  |  |
| OQ-RPT-03 | Dot Matrix Printer Output | User | External dot matrix printer prints the report clearly and completely. |  |  |
| OQ-RPT-04 | Report Export | Authorised User | Reports can be exported in the defined format per role permission. |  |  |
| OQ-RPT-05 | Role-Based Report Access | Administrator/Reviewer/QA/User | Administrator, Reviewer, and QA can view all historical reports; a User can view only their own test reports; no role can edit a finalised report. |  |  |
| OQ-RPT-06 | Historical Report Reprint | Authorised User | A previously approved test report can be retrieved and reprinted on demand; the reprinted copy is legible and complete. |  |  |
| OQ-RPT-07 | Reprinted Report Data Accuracy | Reviewer/QA | On the reprinted historical report, verify the Batch Number, Product Name, Test Start Time, and Test End Time exactly match the original approved record. |  |  |
| OQ-RPT-AT | Audit Trail Check — Reporting | Reviewer/QA | Report generation, export, and reprint actions are recorded in the Audit Trail with user ID, date & time. |  |  |
| 11. AUDIT TRAIL VERIFICATION (CONSOLIDATED) [Ref: 21 CFR 11.10(e)] |  |  |  |  |  |
| OQ-AT-01 | Administrator Audit Trail Access | Administrator | Administrator can view the complete Audit Trail. |  |  |
| OQ-AT-02 | Reviewer Audit Trail Access | Reviewer | Reviewer can view the Audit Trail but cannot modify it. |  |  |
| OQ-AT-03 | QA Audit Trail Access | QA | QA can view the Audit Trail for compliance review. |  |  |
| OQ-AT-04 | User Audit Trail Restriction | User | User cannot access or modify the Audit Trail, unless explicitly permitted by configuration. |  |  |
| OQ-AT-05 | Audit Trail Integrity | System | Audit Trail entries are secure, time-stamped, and non-editable / non-deletable by any user, including Administrator. |  |  |
| OQ-AT-06 | Completeness Check | Reviewer/QA | Every activity performed across Sections 1–10 and 12 of this checklist is confirmed present and accurately recorded in the Audit Trail; no activity is found unrecorded. |  |  |
| 12. ACCESS CONTROL NEGATIVE TESTING [Ref: 21 CFR 11.10(d), 11.10(g)] |  |  |  |  |  |
| OQ-NG-01 | User Attempting Administrator Function | User | System denies access. |  |  |
| OQ-NG-02 | Reviewer Attempting Administrator Function | Reviewer | System denies access beyond Reviewer privileges. |  |  |
| OQ-NG-03 | User Attempting SP/Recipe Approval | User | System denies access; only an authorised approver can approve an SP/Recipe. |  |  |
| OQ-NG-04 | Disabled User Login Attempt | Any Disabled User | Login is blocked with an appropriate alert; the attempt is recorded in the Audit Trail. |  |  |

5. Summary

Overall Compliance Status:      ☐ Compliant      ☐ Non-Compliant      ☐ Compliant with Documented Observations

Observations / Deviations (if any):

6. Sign-off

| Performed By (QC) Name & Signature: ______________________ Date: ______________________ | Reviewed By (Reviewer / QC) Name & Signature: ______________________ Date: ______________________ | Approved By (QA) Name & Signature: ______________________ Date: ______________________ |
| --- | --- | --- |